## Slide 1
PhonoTrack
A Continuous Heart Rate and Blood Pressure Monitoring System from Phonocardiogram
Group No : 03
Shahed Ahmed
Lecturer
Department of Electrical and Electronic Engineering
BUET
Submitted to:
Md. Jahin Alam
Lecturer(PT)
Department of Electrical and Electronic Engineering
BUET
Project: EEE-426

## Slide 2
AGENDA
Phonocardiogram
Crisis in Cardiovascular Health
Existing Solutions
PhonoTrack
Project Overview
Hardware Module
Components
Algorithm Module : Heart Rate
Algorithm Module : Blood Pressure
Dataset Collection
Dataset
Results
Performance Evaluation
Cost Analysis
Future Prospect
Research outcome
Conclusion
Project: EEE-426
02

## Slide 3
Phonocardiogram(PCG)
The sound of life
Figure 01 : Origin of phonocardiogram Signal
Phonocardiogram (PCG) signal represents recording of sounds and murmurs resulting from heart auscultation.
A phonocardiogram (or PCG) is a plot of high-fidelity recording of the sounds and murmurs made by the heart with the help of the machine called the phonocardiograph; thus, phonocardiography is the recording of all the sounds made by the heart during a cardiac cycle.
Aortic pressure, Atrial pressure, Ventricular pressure, Ventricular volume and Electrocardiogram are the origin of PCG signal
Project: EEE-426
[1].
03

## Slide 4
Phonocardiogram(PCG)
The sound of life
Figure 02 : Different label of PCG Signal
Phonocardiogram is an important signal to monitor heart health in a noninvasive way.
PCG captures the two distinct sounds from heart
  S1 : produced by closure of the mitral and tricuspid valves at the beginning of systole.
  S2 : produced by the closure of the aortic and pulmonary valves at the end of systole.
The timing and quality of these sounds can provide valuable information about the heart's function and any potential issues. like-
  Murmurs
  valve disorders,
  congenital heart defects, or
  other structural abnormalities
  Also a sign of heart failure.
Also many important vital signs can be reliably extracted like Heart Rate and Blood Pressure as this data is comparatively cleaner.
Project: EEE-426
04

## Slide 5
Crisis in Cardiovascular Health
Cardiovascular Health Risks in Bangladesh
14.31%
The risk of death due to heart diseases
30%
The risk of death due to heart diseases
2.5 Lakh
Patients die of CVD annually in Bangladesh
[2]
[3]
[4]
[3]
Need for Constant Monitoring
F
A
E
Fitness and exercise
Assessment of medication effectiveness
M
Management of chronic conditions
Early detection of health problems
Project: EEE-426
05

## Slide 6
Existing Solutions
Digital Sphygmomanometer
✅ Reliable
✅ Cost efficient
❌ Continuous measurement
✅ Compact
✅Both Heart rate and BP
ECG
✅ Reliable
❌ Cost efficient
✅ Continuous measurement
❌ Compact
✅ Bothe Heart rate and BP.
Pulse Oximeter
❌ Reliable
✅ Cost efficient
✅ Continuous measurement
✅ Compact
✅ Both Heart rate and BP
Our Proposed system
✅ Reliable
✅ Cost efficient
✅ Continuous measurement
✅ Compact
✅ Both Heart rate and BP
Project: EEE-426
06

## Slide 7
PhonoTrack
A Continuous Monitoring System of HR and BP
Input : Heartbeat Sound
Outputs : (for every 8s)
  Heart Rate
  Blood Pressure
Features:
Continuous Measurement
Powered by charging at regular interval
Weak signal like Bluetooth used, so safe for cardiac patients.
Comparatively cleaner sound acquisition leads to better accuracy.
Cheap and Affordable.
Comfortable and compact for daily usage.
Project: EEE-426
07

## Slide 8
Project Overview
Signal
Acquisition
Data Preprocessing
Sending to
Laptop
Data Framing
Data Denoising
Output Generation
Output Displaying
Hardware Module
Software Module
Project: EEE-426
08

## Slide 9
Hardware Module
Clinical Diaphragm
Microphone Embedded in Diaphragm
PCG Acquisition via Wire/Bluetooth Module
Signal Processing using MATLAB
Output From Signal Processing
Printing in LCD via Arduino
Project: EEE-426
09

## Slide 10
Components
3.5 mm condese
Microphone
Bluetooth Microphone
ECG sensor
Sphygmomanometer
Diaphragm
LCD Display
Arduino uno
Digital Sphygmomanometer
Project: EEE-426
10

## Slide 11
Algorithm Module : Preprocessing
Signal Acquisition via module
Resampling
Normalization
Scaling
Extracting Frames
Denoising
Denoised Signal
Project: EEE-426
11

## Slide 12
Algorithm Module : HR and BP
Denoised Frame
Counting Systole Peaks
Final HR
BP Estimation
Thresholding
Peak Based Downsampling
Feature Extraction
Multivariate Regression
Coefficient Extraction
Semi Empirical Equation for BP
Project: EEE-426
12

## Slide 13
Heart Sound Location
Placement of PhonoTrack device
Project: EEE-426
Area | Aortic | Tricuspid | Mitral | Pulmonary
Location of sound | Second intercostal, right sterna border | Fifth intercostal space, left sternal border | Fifth intercostal space, mid- clavicular border | Second intercostal, left sterna border
Timing of sound | Aortic valve is heard here during S2 | Tricuspid valve is heard here during S1 | Mitral valve is heard here during S1 | Pulmonary valve is heard here during S2
Figure 03: Placement of PhonoTrack device
The region of Tricuspid and pulmonary is the best placement for PCG signal.
13

## Slide 14
Recording System
Setup in Human chest
Diaphragm
Bluetooth
On/Off switch
Recording Device
Belt
System
Dataset Collection
Setup-1 (Bluetooth system)
Project: EEE-426
14

## Slide 15
Recording System
Diaphragm
On/Off switch
Recording Device
Activity: Silent, Not walking
Activity: Normal, walking
Dataset Collection
Setup-2 (Wired system)
Project: EEE-426
15

## Slide 16
Dataset
Figure 04(b): 8s long PCG frame
Figure 04(a): 8s long PCG frame
Our dataset contains mean BP and simultaneous ecg frames for ppg data in 21 subjects with 19 subjects having average HR during data acquisition.
Data was collected from Tricuspid and Pulmonary region to ensure better SNR.
Dataset Descriptions:
Frame Length: 8 Seconds
Average Frames: 15 Frames/subjects
Frame Type: Non-overlapping
Collection region: Tricuspid, Pulmonary
Activities: Standing, sitting, walking,
After light exercise
Sampling Freq: 8000 Hz.
Dataset Summary:
Total PCG Frames : 315 Frames (with ECG + BP)
285 Frames (with avg HR + BP)
Train_test_split : Train dataset (12 subjects)
Test dataset (9 subjects)
Project: EEE-426
16

## Slide 17
Results (Preprocessing Signals)
Figure 05(a) : Original Signal ( 8s frame)
Figure 05(b) : After Resampling at 8000 Hz (8s frame)
Figure 05(c) : After Normalization and Scaling (8s frame)
Project: EEE-426
17

## Slide 18
Results (Frame Extraction)
We have extracted 8s non overlapping frames (a), and for each frame we have used three 4s overlapping frame: 0th -4th sec (b), 2nd -6th sec (c), 4th -8th sec (d) to increase credibility of our Heart Rate estimation.
Project: EEE-426
(a)
(b)
(c)
(d)
18

## Slide 19
Results (After Bandpass filtering)
After framing, Bandpass filtering was done in 20-200 Hz frequency range with 3rd order Butterworth filter to reduce highest frequency noises.
Project: EEE-426
(a)
(b)
(c)
(d)
19

## Slide 20
Results (Frames after wavelet denoising)
In the next step, Wavelet based denoising is used to reduce contribution from other sounds like environmental sounds, breathing sounds etc.
Wavelet based denoising is done in 3 steps:
Wavelet transform to 8 scales
Thresholding each scales.
Reconstructing the frame using inverse wavelet transform.
(a)
(b)
(c)
(c)
Project: EEE-426
20

## Slide 21
Results (HR using Peak Extraction)
For Heart rate calculation we have determined number of peaks in a frame.
The parameters were chosen as follows:
  Minimum delay : 200 ms
  Minimum Peak height : 0.12maxPeak height.
Half of the peaks are taken to consider only the systol peaks from each heartbeat.
Heart rate is then calculated as follows:
HR =
Final heart rate is taken as the average of a non overlapping 8s frame and 3 overlapping 4s frame in that interval by using below equation,
=
Figure 06: predicted Heart rate by HR algorithm
Project: EEE-426
Estimate Heart rate:
21

## Slide 22
Results (BP components extraction)
At first we resampled every 8s frame from 8000 Hz to 2000 Hz, preserving peaks.
After that, we have acquired systole and diastole peaks above a certain range and used these to acquire features.
Project: EEE-426
22

## Slide 23
Results (Semi Empirical Model)
We have acquired total 7 features. They are divided into 2 groups of 4 features for systole and diastole estimation.
Systole Features:
  Systole Decays ()
  Systole Rises ()
  Systole Duration ()
  Heart Rate ()
Diastole features
  Diastole Rises ()
  Diastole Decays ()
  Diastole Duration ()
  Heart Rate ()
Multivariate Linear regression was done on this features to determine our semi empirical equation for Blood Pressure estimation.
= + + + +
= + +
Project: EEE-426
Systole | Values | Diastole | Values
 | 94.558 |  | 65.988
 | -65.286 |  | -26.574
 | 91.504 |  | 48.768
 | 26.217 |  | 22.194
 | 0.199 |  | 0.105
Systole | Values | Diastole | Values
 | 94.558 |  | 65.988
 | -65.286 |  | -26.574
 | 91.504 |  | 48.768
 | 26.217 |  | 22.194
 | 0.199 |  | 0.105
23
Semi empirical model:
Semi empirical model parameter

## Slide 24
Performance Evaluation : HR
Figure 07: Comparison between Label HR and Estimate HR.
24
Project: EEE-426
Figure 08: Error plot of HR in total Dataset with respect to Label..

## Slide 25
Performance Evaluation of BP: Systolic pressure
25
Project: EEE-426
Figure 09: Error plot of BP:Systolic pressure in total Dataset with respect to Label..

## Slide 26
26
Performance Evaluation of BP: Diastolic pressure
Figure 09: Error plot of BP:Systolic pressure in total Dataset with respect to Label..
Project: EEE-426

## Slide 27
Cost Analysis: Project cost
Name of components | cost
Microphone | 1200/-
Bluetooth Microphone | 700/-
Diaphragm(3x) | 900/-
Arduino Microcontroller(2x) | 800/-
Sphygmomanometer | 1500/-
Digital Sphygmomanometer | 3000/-
LCD display | 250/-
ECG sensor | 1200/-
ECG electrode | 1200/-
Miscellaneous | 2000/-
27
Total = TK 11,750/-
Project: EEE-426

## Slide 28
Cost Analysis: Proto-type cost
Name of components | cost
Microphone | 1200/-
Diaphragm(1x) | 300/-
Arduino Microcontroller | 400/-
Rechargeable battery | 300/-
Miscellaneous | 300/-
Total: | Tk. 2400/=
Name of components | cost
Bluetooth Earphone | 700/-
Diaphragm(1x) | 300/-
Arduino Microcontroller | 400/-
Rechargeable battery | 300/-
Miscellaneous | 300/-
Total: | Tk. 2000/=
Setup-01: Wirable device
28
Setup-02: Bluetooth device
Project: EEE-426

## Slide 29
Cost Analysis
JPN1 Omron Digital Blood Pressure Monitor 11000/-
Life Care Blood pressure monitor RAK-288
3000/-
Laser Therapy Watch for Normal Blood Pressure Measurement 20000/-
Manual BP MACHINE
2000 /-
Life Care Blood pressure monitor RAK-288
2700/-
Our proposed device
2000 - 2200/-
29
Project: EEE-426

## Slide 30
Future Prospects
30
This PhonoTrack device will be implement in Automated system by utilizing App Development for Mobile phone.
To Convert MATLAB code into C++ code which is compatible in Microcontroller.
To develop Audio recording module for Microcontroller.
To implement SD card module due to store data in Microcontroller.
Audio recording module
Microcontroller
SD card
module
C++
code
Displays
Project: EEE-426

## Slide 31
Novelty/ Project outcome
31
PhonoTrack device is a continuous monitoring system based on HR and BP.
This project creates new dataset in real-time.
New empirical model for blood pressure was implemented here.
The design of PhonoTrack device was implemented by this project and is very simpler, cheaper, and low maintenance cost than other phonocardiogram device.
PhonoTrack is less noise-sensitive and the SNR of the PCG signal is very high.
Project: EEE-426

## Slide 32
Conclusion
32
PhonoTrack is one kind of Phonocardiography that was implemented in this work and verified with the truth label data.
This project estimated the HR and BP from the PCG signal by using new algorithm and empirical model.
The error of HR is less than 5-10%.
For the BP error is less than 10-15%.
Future research development in sophisticated algorithms and cloud based deep learning artificial monitoring system.
Project: EEE-426

## Slide 33
Thanks, everyone for your patience
32
We are so grateful to…………
Shahed Ahmed
Md. Jahin Alam
Project: EEE-426